## 5E7DBD
***VANES RECALIBRATED AFTER REPAIR 2019/03/01***
S/N: | 611231 | Pair: | 611210
Software download: | www.spingarage.com/611231
Output Format
Type | VCP-over-USB
Baud Rate | 115200
Output Rate | 20 Hz
Output String | $AIRDAQ,FF.FF.FF,ADC,1111,2222,3333,4444<CR><LF>
Element | $AIRDAQ, | FF.FF.FF, | ADC, | 1111, | 2222, | 3333, | 4444 | <CR><LF>
Description | Header | S/N | Output Type | Static Press Counts | Dynamic Press Counts | Alpha Counts | Beta Counts | Carriage Return/Line Feed \x0D\x0A
Vane Calibration Details
Alpha |  |  |  |  | Beta
Tare | 0 | degrees nose up |  |  | Tare | 0 | degrees nose right
| Alpha (deg) | Adj. Deg | Alpha Counts |  |  | Beta (deg) | Adj. Deg | Beta Counts
Nose Up | 30 | 30 | 4864 |  | Nose Left | 30 | 30 | 4436
| 20 | 20 | 4432 |  |  | 20 | 20 | 4045
| 10 | 10 | 3987 |  |  | 10 | 10 | 3658
| 0 | 0 | 3569 |  |  | 0 | 0 | 3259
Nose Down | -10 | -10 | 3134 |  | Nose Right | -10 | -10 | 2862
| -20 | -20 | 2703 |  |  | -20 | -20 | 2490
| -30 | -30 | 2295 |  |  | -30 | -30 | 2094
Linear Best Fit | alpha_deg = 0.0233*alpha_counts -83.1476 |  |  |  | Linear Best Fit | beta_deg = 0.0256*beta_counts -83.5815
Pressure Transducer Details
Static Pressure Transducer |  |  |  |  | Dynamic Pressure Transducer
Make | Honeywell |  |  |  | Make | Honeywell
P/N | HSCMRNN1.6BASA3 |  |  |  | P/N | SSCMRRN100MDSA3
Range | 0-1.6 bar |  |  |  | Range | +/- 100 mbar
Accuracy | 0.25% |  |  |  | Accuracy | 0.25%
Equation (Factory Cal) | ps_bar = 0.00012207*(ps_counts - 1638) |  |  |  | Equation (Factory Cal) | pd_mbar = (0.01525902*(pd_counts - 1638)) - 100
References
https://sensing.honeywell.com/index.php?ci_id=45843
https://sensing.honeywell.com/honeywell-sensing-trustability-hsc-series-high-accuracy-board-mount-pressure-sensors-50099148-a-en.pdf